## Sheet1 (hidden)
Incorrect tank mixing properties | User error | User error | Fails with normal reading | Bump power switch | Hose gets caught on something/someone | Incorrect O2 tank pressure setting | Incorrect O2 tank pressure setting | Bump power switch | Hose gets caught on something/someone | Bump power switch | Tank empty or nearly empty | Tank empty or nearly empty | Assembly error | Pressure sensor malfunction | Pressure sensor malfunction | Unit dropped | Pole clamps not tightened enough | Bolt loosens due to vibration while moving | Unit dropped
320 | 200 | 200 | 180 | 135 | 135 | 120 | 120 | 120 | 120 | 90 | 80 | 80 | 72 | 72 | 72 | 72 | 72 | 72 | 64

## Sheet2 (hidden)
Oxygen Sensor | Air Inlet Fitting | O2 Inlet Fitting | Inspiratory Flow Sensor (Venturi) | Gas Outlet Solenoid | Gas Mixing Tank | Inspiratory Hose Connection | Air Inlet Solenoid | Expiratory Hose Connection | Expiratory Flow Sensor (Venturi)
568 | 432 | 424 | 406 | 271 | 244 | 207 | 197 | 184 | 152

## Sheet3 (hidden)
Incorrect characterization of FiO2 in tank | Sensor Failure | Air supply line connected | Insufficient inlet pressure | O2 supply line connected | Insufficient inlet pressure | Power Failure | Power Failure | Ventilator slides down IV pole | Hose disconnects
320 | 216 | 200 | 200 | 200 | 200 | 162 | 144 | 144 | 135

## FMEA
Product/Process | InVent Ventilator | Date Created | 2020-05-08
Prepared By | Jonathan Schreven & Eric Spaulding
Notes | Last updated on 5/11/2020 by Jonathan Schreven
Function / Requirement | Potential failure mode | Potential effect of failure | SEV | Class | Current process controls | [5 empty] | Recommended actions |  |  | Action results
[5 empty] | Potential cause | Occur | Prevention | Detection | Det | R.P.N. | Recommended actions | Responsible | Target completion date | Actions Taken | Sev | Occur | Det | R.P.N. | Class
O2 Inlet Fitting | Air supply line connected | Incorrect FiO2 delivery to patient | 8 |  | User error | 5 | None | Visual inspection, FiO2 alarm | 5 | 200 | [7 empty] | 0
| External O2 line disconnects | Incorrect/insufficient O2 delivery to patient | 6 |  | Connector breaks | 1 | Locking connector | Visual inspection | 2 | 12 | [7 empty] | 0
| Internal tubing disconnects | Incorrect/insufficient O2 delivery to patient | 6 |  | Tube slips off fitting | 1 | Multi-barbed fitting | Audible gas leak | 2 | 12 | [7 empty] | 0
| Insufficient inlet pressure | Insufficient outlet pressure to drive patient breath | 5 |  | Tank empty or nearly empty | 8 | None | FiO2 alarm | 2 | 80 | [7 empty] | 0
[5 empty] | Incorrect O2 tank pressure setting | 8 | None | FiO2 alarm | 3 | 120 | [7 empty] | 0
[10 empty] | 0 | [7 empty] | 0
O2 Inlet Pressure Regulator | Pressure setting too low | Incorrect FiO2 delivery to patient | 5 |  | Manufacturer quality control issue | 1 | None | FiO2 alarm | 1 | 5 | [7 empty] | 0
| Pressure setting too high | Damage air mixing tank structure | 5 |  | Manufacturer quality control issue | 1 | None | FiO2 Alarm, Audible gas leak | 7 | 35 | [7 empty] | 0
| Tubing disconnect | Incorrect FiO2 delivery to patient | 6 |  | Improper assembly | 2 | None | FiO2 Alarm, Audible gas leak | 2 | 24 | [7 empty] | 0
[10 empty] | 0 | [7 empty] | 0
O2 Inlet Solenoid | Power Failure | N.C. Valve - Oxygen line turns off | 6 |  | Power cord removed from wall/unit | 3 | Cable management | Visual inspection | 1 | 18 | [7 empty] | 0
[5 empty] | Bump power switch | 5 | Flush button on rear of unit | Visual inspection | 3 | 90 | [7 empty] | 0
| Tubing disconnect | Incorrect FiO2 delivery to patient | 6 |  | Improper assembly | 2 | None | Audible gas leak, FiO2 alarm | 2 | 24 | [7 empty] | 0
| Seal fails | Incorrect FiO2 delivery to patient | 5 |  | Manufacturer defect | 1 | None | Audible gas leak, FiO2 alarm | 3 | 15 | [7 empty] | 0
[10 empty] | 0 | [7 empty] | 0
Air Inlet Fitting | O2 supply line connected | Incorrect FiO2 delivery to patient | 8 |  | User error | 5 | None | Visual inspection, FiO2 alarm | 5 | 200 | [7 empty] | 0
| External Air line disconnects | Incorrect/insufficient O2 delivery to patient | 8 |  | Connector breaks | 1 | Locking connector | Visual inspection | 2 | 16 | [7 empty] | 0
| Internal tubing disconnects | Incorrect/insufficient O2 delivery to patient | 8 |  | Tube slips off fitting | 1 | Barbed fitting | Audible gas leak | 2 | 16 | [7 empty] | 0
| Insufficient inlet pressure | Insufficient outlet pressure to drive patient breath | 5 |  | Tank empty or nearly empty | 8 | None | FiO2 alarm | 2 | 80 | [7 empty] | 0
[5 empty] | Incorrect O2 tank pressure setting | 8 | None | FiO2 alarm | 3 | 120 | [7 empty] | 0
[10 empty] | 0 | [7 empty] | 0
Air Inlet Pressure Regulator | Pressure setting too low | Incorrect FiO2 delivery to patient | 5 |  | Manufacturer quality control issue | 1 | None | FiO2 alarm | 1 | 5 | [7 empty] | 0
| Pressure setting too high | Damage air mixing tank | 5 |  | Manufacturer quality control issue | 1 | None | FiO2 Alarm, Audible gas leak | 7 | 35 | [7 empty] | 0
| Tubing disconnect | Incorrect FiO2 delivery to patient | 8 |  | Improper assembly | 2 | None | FiO2 Alarm, Audible gas leak | 2 | 32 | [7 empty] | 0
[10 empty] | 0 | [7 empty] | 0
Air Inlet Solenoid | Power Failure | N.C. Valve - Air line turns off | 8 |  | Power cord removed from wall/unit | 3 | Cable management | Visual inspection | 1 | 24 | [7 empty] | 0
[5 empty] | Bump power switch | 5 | Flush button on rear of unit | Visual inspection | 3 | 120 | [7 empty] | 0
| Tubing disconnect | Incorrect FiO2 delivery to patient | 8 |  | Improper assembly | 2 | None | Audible gas leak, FiO2 alarm | 2 | 32 | [7 empty] | 0
| Seal fails | Incorrect FiO2 delivery to patient | 7 |  | Manufacturer defect | 1 | None | Audible gas leak, FiO2 alarm | 3 | 21 | [7 empty] | 0
[10 empty] | 0 | [7 empty] | 0
Gas Mixing Tank | Hose barb breaks | No gas delivery to patient | 9 |  | Fatigue from hose movement | 3 | Functional prototype cycle testing | Audible gas leak, FiO2 alarm | 2 | 54 | [7 empty] | 0
| Tank damage | No gas to patient | 10 |  | Over pressure | 2 | FEA Analysis, SF = 3 | Audible gas leak, FiO2 alarm | 2 | 40 | [7 empty] | 0
| Tank gasket fails | Insufficient outlet pressure to drive patient breath | 8 |  | Assembly error | 3 | Pressure check at manufacturing | Audible gas leak, FiO2 alarm | 3 | 72 | [7 empty] | 0
[5 empty] | System Overpressure | 1 | Upstream pressure regulators | Audible gas leak, FiO2 alarm | 3 | 24 | [7 empty] | 0
| Tubing disconnect | No gas to patient | 9 |  | Tube slips off fitting | 2 | Barbed fitting | Audible gas leak | 3 | 54 | [7 empty] | 0
[10 empty] | 0 | [7 empty] | 0
Gas Outlet Solenoid | Power Failure | N.C. Valve - No gas to patient | 9 |  | Power cord removed from wall/unit | 3 | Cable management | Visual inspection | 1 | 27 | [7 empty] | 0
[5 empty] | Bump power switch | 5 | Flush button on rear of unit | Visual inspection | 3 | 135 | [7 empty] | 0
| Tubing disconnect | Incorrect FiO2 delivery to patient | 9 |  | Improper assembly | 2 | None | Audible gas leak, FiO2 alarm | 2 | 36 | [7 empty] | 0
| Seal fails | Incorrect FiO2 delivery to patient | 7 |  | Manufacturer defect | 1 | None | Audible gas leak, FiO2 alarm | 3 | 21 | [7 empty] | 0
| Valve sticks open | Excessive pressure delivered to patient | 9 |  | Mechanical valve failure | 2 | Emergency pressure relief valve | PIP monitor | 1 | 18 | [7 empty] | 0
[5 empty] | Erroneous control system signal | 2 | System testing, emergency pressure relief valve | PIP monitor | 1 | 18 | [7 empty] | 0
| Valve doesn't open enough | Insufficient air volume/flow rate to patient | 8 |  | Erroneous control system signal | 2 | System testing | Tidal Volume not achieved alarm | 1 | 16 | [7 empty] | 0
[10 empty] | 0 | [7 empty] | 0
Inspiratory Flow Sensor (Venturi) | Sensor tubing disconnects | Incorrect gas delivery to patient | 9 |  | Improper assembly | 2 | Pressure check at manufacturing | Tidal Volume not achieved alarm | 3 | 54 | System pause for suspected sensor failure | [6 empty] | 0
| Hose disconnects | No gas delivery to patient | 9 |  | Tube slips off fitting | 2 | Barbed fitting | Audible gas leak | 3 | 54 | [7 empty] | 0
| Inlet hose barb breaks | No gas delivery to patient | 9 |  | Fatigue from hose movement | 3 | Functional prototype cycle testing | Audible gas leak, FiO2 alarm | 2 | 54 | [7 empty] | 0
| Sensor barb breaks | Incorrect tidal volume to patient | 10 |  | Fatigue from hose movement | 3 | Functional prototype cycle testing | Audible gas leak, FiO2 alarm | 2 | 60 | NPT barbs? | [6 empty] | 0
| Incorrect flow measure | Incorrect tidal volume to patient | 10 |  | Differential pressure sensor malfunction | 2 | Testing against calibrated flow sensor | Tidal volume not achieved alarm | 2 | 40 | [7 empty] | 0
| Incorrect pressure measurement | Excess pressure to patient | 9 |  | Pressure sensor malfunction | 2 | None | Insp. Pressure not achieved alarm | 4 | 72 | [7 empty] | 0
| Incorrect pressure measurement | Insufficient pressure to patient | 9 |  | Pressure sensor malfunction | 2 | None | Insp. Pressure not achieved alarm | 4 | 72 | [7 empty] | 0
[10 empty] | 0 | [7 empty] | 0
Inspiratory Hose Connection | Hose disconnects | No delivery to paitient | 9 |  | Hose gets caught on something/someone | 3 | Designed to ISO standard, vertical port orientation | External patient sensors (e.g. pulse ox.) | 5 | 135 | [7 empty] | 0
| Connection breaks | Incorrect patient gas delivery | 9 |  | Unit dropped | 2 | Designed to ISO standard, port protected behind display | Insp. Pressure and Tidal volume not achieved alarms | 4 | 72 | [7 empty] | 0
[10 empty] | 0 | [7 empty] | 0
Expiratory Hose Connection | Hose disconnects | Incorrect expiratory pressure | 8 |  | Hose gets caught on something/someone | 3 | Designed to ISO standard, vertical port orientation | External patient sensors (e.g. pulse ox.) | 5 | 120 | [7 empty] | 0
| Connection breaks | Incorrect patient gas delivery | 8 |  | Unit dropped | 2 | Designed to ISO standard, port protected behind display | Insp. Pressure and Tidal volume not achieved alarms | 4 | 64 | [7 empty] | 0
[10 empty] | 0 | [7 empty] | 0
Expiratory Flow Sensor (Venturi) | Sensor tubing disconnects | Sensors produce false data | 8 |  | Improper assembly | 2 | Pressure check at manufacturing | Tidal Volume not achieved alarm | 3 | 48 | [7 empty] | 0
| Hose disconnects | Unable to measure expiratory flow and pressure | 8 |  | Tube slips off fitting | 2 | Barbed fitting | Audible gas leak | 3 | 48 | [7 empty] | 0
| Incorrect flow measure | Incorrect measure of patient breating status | 6 |  | Differential pressure sensor malfunction | 2 | Testing against calibrated flow sensor | Tidal volume not achieved alarm | 2 | 24 | [7 empty] | 0
| Incorrect PEEP measurement | Incorrect data output | 4 |  | Pressure sensor malfunction | 2 | None | Insp. Pressure not achieved alarm | 4 | 32 | [7 empty] | 0
[10 empty] | 0 | [7 empty] | 0
Unit Mounting | Ventilator slides down IV pole | Patient circuit and/or power disconnect | 9 |  | Pole clamps not tightened enough | 2 | None | Visual observation | 4 | 72 | [7 empty] | 0
[5 empty] | Bolt loosens due to vibration while moving | 2 | None | Visual observation | 4 | 72 | [7 empty] | 0
[10 empty] | 0 | [7 empty] | 0
Oxygen Sensor | Sensor Failure | Incorrect FiO2 delivery to patient | 9 |  | Fails with normal reading | 2 | None | None | 10 | 180 | [7 empty] | 0
[5 empty] | Fails with extreme value | 2 | None | Insp. Pressure not achieved or overpressure alarm | 2 | 36 | [7 empty] | 0
| Oring failure | Insufficient outlet pressure to drive patient breath | 8 |  | Over compression of o-ring | 2 | None | Audible gas leak, tidal volume and insp. Pressure not achieved alarms | 2 | 32 | [7 empty] | 0
| Incorrect characterization of FiO2 in tank | Incorrect FiO2 delivery to patient | 10 |  | Incorrect tank mixing properties | 4 | Testing against calibrated O2 sensor | Visual observation of patient by care taker | 8 | 320 | [7 empty] | 0
[10 empty] | 0 | [7 empty] | 0
[10 empty] | 0 | [7 empty] | 0
[10 empty] | 0 | [7 empty] | 0
[10 empty] | 0 | [7 empty] | 0
[10 empty] | 0 | [7 empty] | 0
[10 empty] | 0 | [7 empty] | 0
[10 empty] | 0 | [7 empty] | 0
[10 empty] | 0 | [7 empty] | 0
[10 empty] | 0 | [7 empty] | 0
[10 empty] | 0 | [7 empty] | 0
[10 empty] | 0 | [7 empty] | 0
[10 empty] | 0 | [7 empty] | 0
[10 empty] | 0 | [7 empty] | 0
[10 empty] | 0 | [7 empty] | 0
[10 empty] | 0 | [7 empty] | 0
[10 empty] | 0 | [7 empty] | 0
[10 empty] | 0 | [7 empty] | 0
[10 empty] | 0 | [7 empty] | 0
[10 empty] | 0 | [7 empty] | 0
[10 empty] | 0 | [7 empty] | 0
[10 empty] | 0 | [7 empty] | 0
FMEA Tips Modifying To add, insert, or remove rows or columns, right click and select "Add to FMEA", "Insert Into FMEA", or "Remove from FMEA". Please do not use Excel's Insert or Delete command to add rows or columns. Functions, Potential Failure Modes, and Effects Identify the functions of the design/product/process and enter into the "Function / Requirement" column. For each function or requirement, brainstorm the possible failure modes. For each failure mode, identify the possible effect(s) and place into the column "Potential effect of failure". Numerical Analysis SEV = Severity (1 = Least Severe, 10 = Most Severe) Occur = Probability of Occurance (1 = Least Likely, 10 = Most Likely) Det = Probability of Detection (1 = Most Likely, 10 = Least Likely) For more information, select "Quantum XL" > "About (Help)…" and click on the "Help" button, or visit us at www.sigmazone.com. Copyright(C) 2008-2013 SigmaZone

## FMEA Ratings
FMEA Rating Tables |  | These definitions have been provided as a reference only, your industry may use different Severity, Occurrence, or Detection values. A list of other organizations with links to their ratings is at the bottom of this sheet
Severity of Effect |  |  |  |  | Likelihood of Detection
10 | May result in safety issue or regulatory violation without warning |  | Safety / Regulatory / Legal Zone |  | 10 | Absolutely uncertain that failure will be detected
9 | May result in safety issue or regulatory violation with warning |  |  |  | 9 | Very remote chance that failure will be detected
8 | Primary function is lost or seriously degraded |  | Warranty/ Field Failure Zone |  | 8 | Remote chance that failure will be detected
7 | Primary function is reduced and customer is impacted |  |  |  | 7 | Very low chance that failure will be detected
6 | Secondary function is lost or seriously degraded |  |  |  | 6 | Low chance that failure will be detected
5 | Secondary function is reduced and customer is impacted |  |  |  | 5 | Moderate chance that failure will be detected
4 | Loss of function or appearance such that most customers would return product or stop using service |  |  |  | 4 | Moderately high chance that failure will be detected
3 | Loss of function or appearance that is noticed by customers but would not result in a return or loss of service |  |  |  | 3 | High chance that failure will be detected
2 | Loss of function or appearance that is unlinkely to be noticed by customers and would not result in a return or loss of service |  |  |  | 2 | Very high chance that failure will be detected
1 | Little to no impact |  |  |  | 1 | Almost certainty that failure will be detected
[5 empty] | Class Column Symbol Table | [5 empty] | SRL | Safety/Regulatory/Legal
[11 empty] | WFF | Warranty/Field Failure
Probability of Occurrence |  |  |  | Severity | 10 |  | SRL | SRL | SRL | SRL | SRL | SRL | SRL | SRL | SRL
10 | 1 in 2 |  |  |  | 9 |  | SRL | SRL | SRL | SRL | SRL | SRL | SRL | SRL | SRL
9 | 1 in 10 |  |  |  | 8 |  |  |  | WFF | WFF | WFF | WFF | WFF | WFF | WFF
8 | 1 in 50 |  |  |  | 7 |  |  |  | WFF | WFF | WFF | WFF | WFF | WFF | WFF
7 | 1 in 250 |  |  |  | 6 |  |  |  | WFF | WFF | WFF | WFF | WFF | WFF | WFF
6 | 1 in 1,000 |  |  |  | 5 |  |  |  | WFF | WFF | WFF | WFF | WFF | WFF | WFF
5 | 1 in 5,000 |  |  |  | 4 |  |  |  | WFF | WFF | WFF | WFF | WFF | WFF | WFF
4 | 1 in 10,000 |  |  |  | 3
3 | 1 in 50,000 |  |  |  | 2
2 | 1 in 250,000 |  |  |  | 1
1 | 1 in 1 Million |  |  |  |  | 1 | 2 | 3 | 4 | 5 | 6 | 7 | 8 | 9 | 10
Occurrence
These definitions have been provided as a reference only, your industry may use different Severity, Occurrence, or Detection values. Each of the following organizations has their own standard:
SAE J1739: http://www.sae.org/technical/standards/J1739_200208 SAE ARP5580: http://www.sae.org/technical/standards/ARP5580 AIAG: https://www.aiag.org ISO 9000: http://www.iso.org/iso/iso_catalogue/catalogue_tc/catalogue_detail.htm?csnumber=42180 ISO/TS 16949:2002: http://www.iso.org/iso/catalogue_detail?csnumber=36155